## 100-262 Rev. D- Standards Set 1
|  | Instrument ID: |  |  | B3 Film Peak Testing
|  | Technician Name: |  |  | All testing is performed with the instrument in the Absorbance mode
|  | Date: |  |  | Test Wavelength (nm) | Absorbance Result
|  | Temperature (C or F): |  |  | 545 |  | Dosimeter ID#:
|  | Calibration Standards Set ID#: |  |  | 546
|  | Calibration Certificate #: |  |  | 547
548
| Spectronic Standards Calibration Verification Tests for Genesys 20 |  |  |  | 549
| All testing is performed with the instrument in the Transmittance mode |  |  |  | 550
551
| 0% Transmittance and Stray Radiant Energy Test |  |  |  | 552 |  | Expected Peak(nm): | 552
|  | %T Result | Acceptance Criteria | Pass/Fail | 553 |  | Peak Found at(nm):
| 0% T at 400nm: | Enter | 0 |  | 554 |  | Acceptance Criteria: | ±2nm
| SRE at 400nm: | Enter | ≤0.45 |  | 555
| SRE at 340nm: | Enter | ≤0.45 |  | 556 |  | Pass/Fail
557
| Photometric Performance |  |  |  | 558
| Certified Value at 590nm | %T Result | Acceptance Criteria | Pass/Fail | 559
(10% T) |  |  | ±1.2 %T
(50% T) |  |  | ±1.2 %T
| 1st Wavelength Accuracy Test |  |  |  | 2nd Wavelength Accuracy Test
| Certified Wavelength (nm): | %T Result |  |  | Certified Wavelength (nm): | %T Result
| -7 |  |  |  | -7
| -6 |  |  |  | -6
| -5 |  |  |  | -5
| -4 |  |  |  | -4
| -3 |  |  |  | -3
| -2 |  | Standard Peak Value(nm): | 0 | -2 |  | Standard Peak Value(nm): | 0
| -1 |  | Peak Found at(nm): |  | -1 |  | Peak Found at(nm):
|  |  | Acceptance Criteria: | ±3nm |  |  | Acceptance Criteria: | ±3nm
| 1 |  |  |  | 1
| 2 |  | Pass/Fail |  | 2 |  | Pass/Fail
| 3 |  |  |  | 3
| 4 |  | *Use certified wavelength value with SBW closest to 8nm spectral bandwidth |  | 4 |  | *Use certified wavelength value with SBW closest to 8nm spectral bandwidth
| 5 |  |  |  | 5
| 6 |  |  |  | 6
| 7 |  |  |  | 7
|  | Signature: |  |  | Date:
|  | Reviewed By: |  |  | Date:

## 100-262 Rev. D- Standards Set 2
Instrument ID:
Technician Name:
|  | Date: |  |  | B3 Film Peak Testing
|  | Temperature (C or F): |  |  | All testing is performed with the instrument in the Absorbance mode
|  | Calibration Standards Set ID#: |  |  | Test Wavelength (nm) | Absorbance Result
|  | Calibration Certificate #: |  |  | 545 |  | Dosimeter ID#:
546
| Spectronic Standards Calibration Verification Tests for Genesys 20 |  |  |  | 547
| All testing is performed with the instrument in the Transmittance mode |  |  |  | 548
549
| 0% Transmittance and Stray Radiant Energy Test |  |  |  | 550
|  | %T Result | Acceptance Criteria | Pass/Fail | 551
| 0% T at 400nm: | Enter | 0 |  | 552 |  | Expected Peak(nm): | 552
| SRE at 400nm: | Enter | ≤0.45 |  | 553 |  | Peak Found at(nm):
[5 empty] | 554 |  | Acceptance Criteria: | ±2nm
| Photometric Performance |  |  |  | 555
| Certified Value at 590nm | %T Result | Acceptance Criteria | Pass/Fail | 556 |  | Pass/Fail
(3% T) |  |  | ±1.2 %T |  | 557
(10% T) |  |  | ±1.2 %T |  | 558
(30% T) |  |  | ±1.2 %T |  | 559
(50% T) |  |  | ±1.2 %T
| 1st Wavelength Accuracy Test |  |  |  | 2nd Wavelength Accuracy Test
| Certified Wavelength (nm): | %T Result |  |  | Certified Wavelength (nm): | %T Result
| -7 |  |  |  | -7
| -6 |  |  |  | -6
| -5 |  |  |  | -5
| -4 |  |  |  | -4
| -3 |  |  |  | -3
| -2 |  | Standard Peak Value(nm): | 0 | -2 |  | Standard Peak Value(nm): | 0
| -1 |  | Peak Found at(nm): |  | -1 |  | Peak Found at(nm):
|  |  | Acceptance Criteria: | ±3nm |  |  | Acceptance Criteria: | ±3nm
| 1 |  |  |  | 1
| 2 |  | Pass/Fail |  | 2 |  | Pass/Fail
| 3 |  |  |  | 3
| 4 |  | *Use certified wavelength value with SBW closest to 8nm spectral bandwidth |  | 4 |  | *Use certified wavelength value with SBW closest to 8nm spectral bandwidth
| 5 |  |  |  | 5
| 6 |  |  |  | 6
| 7 |  |  |  | 7
|  | Signature: |  |  | Date:
|  | Reviewed By: |  |  | Date: